Request Title:  	Use of Parylene-D in Parcoater

Requestor: 		Paul Comita					Date:  Mar 4, 2020

Non-SNF staff, please complete the following:

Badger ID:	pcomita50@gmail.com

Company name/PI: Silicon Microstructures/ Fernando Alfaro

Send completed form to: snf-promcommittee@lists.stanford.edu

- The chemical or material. trade names, and CAS numbers where appropriate.

Common name: Parylene-D

Trade name: DPX-D

CAS name: Tricyclo[8.2.2.2(sup 4,7)]hexadeca-4,6,10,12,13,15-hexaene, tetrachloro-

CAS Number: 	30501-29-2

See attached MSDS at end of file.  Read the MSDSs as well as the Stanford Chemical Storage Groups to determine the Storage Group Identifier and Main Hazard Class of your chemical/material.

- Vendor/manufacturer info:

https://scscoatings.com/parylene-coatings/

Specialty Coating Systems +1-317-451-8549

- Reason for request:

Parylene-D is a higher temperature form of Parylene, so it is used as an alternative to Parylene-N or C for higher temperature applications.    Parylene D, the third member of the series, is produced from the same raw material as Parylene N, modified by the substitution of chlorine atoms for two of the aromatic hydrogens. Parylene D is similar in properties to Parylene C with the added ability to withstand slightly higher use temperatures.

- Process Flow: Please provide a detailed process flow description on how and where you proposed to use this chemical. This should include all Lab equipment to be used for processing your wafers once your new chemical or material has been used (even if your new material is a film that is removed, it may still pose potential contamination concerns.)

The Parylene D will be used in the same lab as presently is the Parylene C used, Allen ,

Room 155. Use different measuring utensils in order NOT to mix or contaminated the 2

compositions. Clean the working bench before and after use.

The differences between the parylene C and D are :

Process temperature C @ 690C vs. D processed at    700C

Vaporizer temperature : C at 175C and D @ 175C

Vacuum for C @ 26 torr, vs D @ 20torr

Baffle: turn the deposition chamber's tubular baffle 180 degrees so the holes face the middle of the chamber

- Amount and form. How much will you bring in? Is it solid, powder or liquid? (Note: as a general rule, powders are not permitted in the cleanroom.) Do you need to mix it to use it?

No mix is needed. Keep it in original factory container

- Storage: Will you be storing your chemical/material at SNF? If so, please note any potential reactivity’s (this should be on the MSDS). Storage groups A, B, D and L are stored in the yellow solvent cabinet in the furnace support area, while storage groups C, E, F and G are stored on top of one of the Pass-through Carts. Ensure your chemical container or material is properly labeled. If there is no available room, it must be stored by in the bulk storage area. You will then need to obtain it from receiving area personnel each time you want to use it and return it to them when you are finished using it (or each time you leave the lab). Note that there is no storage of chemicals/materials in the processing lab or at any wet bench.

Stored in one of the bench designated by the SNF staff, along with the existing Parylene C and N.

- Disposal: How will you dispose of any waste or excess chemical or material? In your discussions with experts and vendors, try to determine the best way to dispose of your spent chemicals and by-products. Please refer to the SNF Lab members Safety Manual for the different methods of waste disposal that are available in the lab.

No residue will be generated, all processed during the reaction in the chamber.

To be completed by the PROM COMMITTEE:

Special handling requirements:

On-going approval or staff interaction requirements:

Approved By/Date: